Interface Control Document

of the LCGT Project

| Subgroup Name | Data Acquisition System |
| --- | --- |
| Subgroup Leader | Daisuke Tatsumi National astronomical observatory of Japan |

Glossary

| GPS | Global Positioning system |
| --- | --- |
| DAQ | Data Acquisition System |
| ADC | Analog to Digital Converter |
| bps | bit per second (a unit to indicate data flow capability) |

1. Definition of subsystem

This subsystem should collect data from an interferometric detector.

Data storage, distribution and pre-procession for GW search are also provided for collaborators.

2. Task area

| No. | Function | Hardware, software |
| --- | --- | --- |
| 1 | Main DAQ system | Analog Digital Converter, Signal conditioning filter (whitening filter) |
| 2 | Frame Maker | Computer: Real time operation system to storage data with GW FRAME FORMAT. |
| 3 | Detector Diagnosis | Computer: Online data analysis soft for detector diagnosis and trend-data archive |
| 4 | Environment Monitor | Computer: Environmental data monitoring soft, display for user interface. |
| 5 | Time Keeper and Global Positioning System | GPS, GPS signal transmission system, GPS time server with Rubidium atomic clock, sampling clock distribution system for ADC and DAC. |
| 6 | Raw Data Archive | Computer, tape archive system and data storage system: Archived data database and its human interface. |
| 7 | Pre Analysis Server | Computer: Data re-formatting for GW signal search |
| 8 | Data Distribution | Computer and high-speed network: Data distribution system for collaborator's demands |

2. Required specifications

|  | Item | Requirements |
| --- | --- | --- |
| Main DAQ | Sampling rate | 32768 Hz (down-sample from 65536 Hz) |
|  | Number of channels | 32 |
|  | Bit resolution | 24 bit (4 byte integer) |
|  | Anti-alias filter | 10 kHz cut-off |
|  | Data rate | 4 MB/sec |
|  | Frame length | 1 second (32768 samples) |
| Frame builder | Operation system | UNIX |
|  | (minimum) Hard disk | 14.4 GB/hour x 70 hours = 1 TB |
|  | Network interface | Gigabit Ethernet with optical fiber |
| Detector diagnosis | Sampling rate | 16384 Hz (down-sample from 65536 Hz) |
|  | Number of channels | 512 CH = 64 CH x 8 stations |
|  | Bit resolution | 16 bit (2 byte integer) |
|  | Anti-alias filter | 5 kHz cut-off |
|  | Data rate | 2 MB/sec x 8 stations |
| Environmental monitor | Sampling rate | 32 Hz (down-sample from 65536 Hz) |
|  | Number of channels | 512 CH = 64 CH x 8 stations |
|  | Bit resolution | 16 bit (2 byte integer) |
|  | Anti-alias filter | 10 Hz cut-off |
|  | Data rate | 2 kB/sec x 8 stations |
| Timing system and Global Positioning System | Time accuracy | < 1 micro sec |
| Raw data archive | from Main DAQ | 4 MB/sec | = 14.1 GB/hour |
|  | from Diagnosis | 2 MB/sec x 8 | = 56.3 GB/hour |
|  | from Environmental | 2 kB/sec x 8 | = 0.05 GB/hour |
|  | Total data rate | 20 MB/sec | = 70.4 GB/hour |
|  | Total data storage | 100 TB in minimum |
| Pre analysis server | Computing power | 300G Flops in minimum |
| Data Distribution | Network interface | Gigabit Ethernet with optical fiber |

3. Interface between systems

| Item | Related sub-system | Requirement | Confirmation |
| --- | --- | --- | --- |
| Observation band | Data analysis | 10 Hz - 10 kHz | by DAS |
| ADC bit resolution | Data analysis | > 16 bits | by DAS |
| Number of channels for GW search | Data analysis | 32 channels | by DAS |
| Pre-processed data formats | Data analysis | TBD | not yet |
| Acceptable delay time to distribute observation data for collaborators | Data analysis | < 1 hour | not yet |
| Data storage capability | Data analysis | 100T bytes in minimum | not yet |
| Accuracy of time | Data analysis | < 1us | not yet |
| Computational capability for GW searches | Data analysis | 300G Flops in minimum | not yet |
| Network capability | Infrastructure | 10Gbps spec network switches Each 16 ports for 8 DAQ stations | not yet |
| Power consumption | Infrastructure | 100 kW | not yet |
| Waist heat | Infrastructure | 90 kW | not yet |
| Location and space | Infrastructure | 24 m2 + 8 Racks | not yet |
| Installation schedule | Management | until 4th year end | not yet |
| Budget allocation | Management | 50,000,000 Jpn Yen | not yet |
| Man power allocation | Management | 2 person x 5 years | not yet |