Current Status of the LCGT Interferometer Design
Yoichi Aso on behalf of the LCGT ISC Group
2010/9/27 LCGT F2F Meeting

---
Summary of the current status
Working to fix parameters necessary for procurements
Parameters relevant to tunnel digging
Lengths of various parts of the IFO
Central part topology, i.e. folding or not
Parameters relevant for mirror order
Size, reflectivity, ROC, wedge, error tolerance, etc
TO DO
Detailed optical layout
Alignment Control
Green lock design
Double-check of the design (review)

---
Basic Interferometer Parameters
Arm length	3km
Arm Cavity Finesse	1550
Arm Cavity g-factor	sqrt(1/3) / or -sqrt(1/3)	Negative g-factor reduces the optical spring instability
PRC length	73.3m	Folded
SRC length	73.3m	Folded
MI asymmetry	3.33m
PRG	10
SRG	24
Test Mass Size	Φ25cm x t15cm	Sapphire
BS Size	Φ ~ 40cm
PRM, SRM size	Φ >= 10cm

---
Mirror name conventions
ETMY
ITMY
PRM
PR2
ITMX
ETMX
BS
PR3
SR2
SR3
SRM

---
Arm Cavity g-factor
Factors to consider
Cavity must be stable (no low-order HOMs should not resonate)
Angular instability caused by radiation pressure should be small enough
The parametric instability also depends on the g-factor
sqrt(1/3) --> L=3000m, ROC=7098.08m (Conventional number)
-sqrt(1/3) --> L=3000m, ROC=1901.92m (Alternative)
g-factor
Candidates
Beam Size: 3.5cm for both cases
HOM degeneration: the same for both cases
Angular instability:
Eigen-frequencies: 1.66Hz and 0.86Hz
Positive g-factor 1.66Hz is unstable
Negative g-factor 0.86Hz is unstable
Different WFS bandwidth
Parametric Instability:
There are preferred regions in the g factor space from the view point of PI.
The regions are the same for positive and negative g.
However, dg/dR is different by a factor of 13 (negative g is larger).
Therefore, negative g is more sensitive to the ROC error (not limited
to the PI actually).
Negative g-factor is more prone to the mirror ROC error for PI

---
PI calculation by K. Yamamoto
Negative g-factor
Positive g-factor
The error requirement on the mirror ROC is stricter for the negative g-factor.

---
Folding
Why folding is necessary ?
Degenerated recycling cavities: Source of headaches
(Sideband mode is not determined, unstable control signal, bad contrast, GW sideband reduced)
We want to stabilize the RCs, but how much ? (to be assessed by MIST and Finesse)
According to LIGO and Virgo studies, one way Gouy phase shift of 20deg seems reasonable
Is folding possible within the LCGT constraints ?
Yes, there are solutions
Many factors to consider:
Beam spot size, susceptibility to errors, ease of adjustment, back scattering, etc ...
Uniquely determined by the geometric requirement to
minimize the incident angles (to avoid astigmatism)
Folding Mirror Positions
Necessary for vacuum system design

---
Folding configuration by D. Tatsumi

---
Effect of the ROC errors in the folding mirrors
We have to be able to adjust the folding mirror positions by +/-5cm or so.

---
Mirror Size
Factors to consider
It is desirable to have as few as possible variations of size and weight of the mirrors
from the view point of SAS
Especially, changing the weight is costly.
What is the maximum possible mirror size ?
Mirror Size Requirements from IFO's point of view

---
Optical Loss in the PRC by BS size effect by H. Yamamoto

---
Mirror Size Plans
Plan A: 16kg plan
For iLCGT: All the mirrors are 16kg fused silica
For bLCGT: ITMs and ETMs are replaced with 30kg Sapphire ones
Suspension
For iLCGT, ITMs and ETMs are suspended by the Type B platform
It will be replaced by Type A SAS in bLCGT
Everything else can be suspended by Type B SAS
Plan B: 30kg plan
For iLCGT:
ITM, ETM, BS, PR3, SR3: 38cmx12cm 30kg fused silica mirrors
PRM, PR2, SRM, SR2: 10cm TAMA sized mirrors
For bLCGT: ITMs and ETMs are 25cmx15cm 30kg sapphire mirrors
Suspension
No need to change suspensions between iLCGT and bLCGT
There will be more TypeA SAS
10cm optics could be suspended with simpler suspensions

---
Optical Layout Tool
A python module to automatically compute the propagation, reflection,
 deflection of Gaussian beams.
Programmatic design of the IFO optical layout to satisfy various quantitative
 requirements at the same time

---
Power Recycling Cavity Folding Part
PRM
PR2
BS
PR3
SR2

---
Tunnel Slope
LCGT Tunnels slope by 1/200
for water drainage
Mirrors are tilted
Vertical-to-Horizontal coupling > 1/200
The slope direction (high-to-low): X-end -> Center -> Y-end
BS must also be tilted
Seems like this is unavoidable
Where do we impose the slope ?
Only the arms ?
Entire interferometer ?
We have no experience with an interferometer not on a single plane
Put the entire interferometer on a single (but slanted) plane

---
Lock Acquisition
Basic ideas for green laser lock
Green Lasers are injected from PR3 for X-arm, and SR3 for Y-arm
PR3, SR3 and BS are transparent to green.
ITMs and ETMs are dichroic mirrors
Each green laser is locked to the main laser through a PLL
Additional AOM could be inserted between the green lasers and PLL
 to provide fast frequency actuation
Main Laser
PRM
PR2
AOM
Green1
PLL1
PR3
SR2
PDH1
Lock procedure
Each arm is locked to the green laser by
 frequency feedback to the PLLs and AOMs.
Common signal of the two arms is fed back
 to the main laser (high freq.) and CARM (low freq.)
Differential signal is fed back to DARM
Scan the PLL1 and PLL2 to find the resonance
 of the main laser
SR3
SRM
PDH2
PLL2
Green2
AOM

---
Summary and TO DO List
Parameters necessary for Tunnel, Vacuum and Mirrors are almost fixed
Mirror specs needs to be discussed more.
But sizes are soon to be fixed.
TO DO
Find an ASC Scheme
Check the GW signal degradation with MIST and Finesse
Finalize folding design
Finalize the green locking design
Finalize mirror specs
Finalize the optical layout
Determine the isolation requirements for RC suspensions
etc ...