## Slide 1: LCGT Vacuum System
Items
  ** beam tubes of 3-km long and 800 mm in diameter
  ** optical baffles distributed in the beam tubes
  ** chambers for the mirrors with suspension and vibration-isolation devices
  ** vacuum pumping system
  ** overall layout
120123 VAC (YS)
Interface
  all of the subsystems
    ** choice of materials for components used in cryostat, optical devices, vibration-isolation devices, , …
    ** interferometer layout and operation sequence
cf. vacuum system philosophy in accelerator
  ** beam dynamics and vacuum system have to be designed simultaneously.
  ** choice of materials having low outgassing is most important.
  ** mechanical and electric design has to be simple.
  ** materials properties have to be examined and measured before choosing.
(KEK Internal 07-17)
1. Overview

## Slide 2
  ** We predict the noise dx due to the residual gas (water) molecules;
  for dx = 1×10-21 m/√Hz@100 Hz (safety margin of 10),
  the pressure in the beam tubes is to be kept at 2×10-7 Pa.
      (h=3×10-24 /(Hz)1/2 @100Hz corresponds to dx=1×10-20 m/(Hz)1/2)
<<scattering effect due to residual gas>>
120123 VAC (YS)
  ** For long-term and stable operation of the interferometer, the vacuum system is to be designed so as to minimize the maintenance work, and so as to shorten the pump-down time. A long lifetime of the vacuum component and low outgassing is required.
<<minimize the maintenance work>>
Ti cathode
Penning cell
Required Pressure：in the order of 10-7 Pa, or lower
LCGT Vacuum System
2. System design

## Slide 3
LCGT Vacuum System
120123 VAC (YS)
ETMYA
ITMY
ETMXA
ITMX
IP
TMP
FL P
BS
PRM
PR3
PR2
SRM
SR3
SR2
PD
MCF
MC
3000 m
double chambers (2.4 and 1.5 m in dia.)
//GASF + I-Pendulum + cryogenic//
chambers (1.5 m in dia./2 m for BS)
//GASF + I-Pendulum//
chambers (2 m in dia.)
//stack + D-Pendulum//
iLCGT
iLCGT
MMT
ETMYB
ETMXB
less changes “i” to “b”
2. System design

## Slide 4
** “surface passivation” is performed by electro-polish followed by bake.
    expected outgassing rate is on the order of 10-8 Pa m3 m-2 s-1, or lower.
    Rmax 3 mm, Ra 0.5 mm.
** “flange connection” with metal O-ring gasket (silver plated) is chosen.
    humidity test of the gasket shows erosion proof.
** tubes of “mirror finish by Electro-Chemical Buffing” is to be installed in the mid region of 800-m long. Rmax 0.2 mm, Ra 0.03 mm
120123 VAC (YS)
Electro-polished tube of 12 m long
LCGT Vacuum System
beam tube (478 of 12-m long and 0.8 m in diameter)
2. System design

## Slide 5
N=(h/d)/(L/R)
** estimated scattered light noise at Kamioka
  dx = 3×10-21 m/√Hz@30Hz (tube vibration amplitude=1×10-11 m/√Hz assumed )
** For more margin, baffles at “every 12 meters”. dx = 5×10-22 m/√Hz
** For randomizing phase of edge-scattered light, baffles with “saw-tooth edge”.
120123 VAC (YS)
LCGT Vacuum System
optical baffle (every12-m along the arm, 40-mm in height, 45-degree tilted)
2. System design

## Slide 6
  ** measured outgassing rate of DLC is 4×10-9 Pa m3 m-2 s-1,
120123 VAC (YS)
真空
LCGT Vacuum System
optical baffle (diamond-like-carbon/DLC coated)
2. System design

## Slide 7
  ** 17 of 21 chambers are operated at room temperature, 4 for cryogenic system
  ** installed materials of elastomer and plastomer should be investigated.
** Although the aluminum-coated thin PET (polyethylene terephtalate) film is suitable material for thermal shield,
the outgassing rate is higher than those of metal surfaces.
** Outgassing rate of a PET film of 12 micrometer thick is measured.
The rate decreases to 10-6 Pa m3 s-1 m-2 for about 10h, then reaching to the order of 10-8 Pa m3 s-1 m-2 for 200h.
** Water molecules absorbed in film is possibly diffused to the surface
and desorbed with a long period of 100 hours.
120123 VAC (YS)
LCGT Vacuum System
chambers
2. System design

## Slide 8
** pumping unit consisted of dry-pump, TMP and ion-pump is distributed “every 100 meters” along tubes.
ETMX
ITMX
IP
TMP
FL P
3000 m
** expected pump-down scheme (a 3-km arm)
  to 1 Pa >> few days by dry-pump
  to 10-6 Pa >> 50 hours by TMP
  to 10-7 Pa >> 500 hours by IP
(based on the ougassing rate in test tubes)
** pumping speed of the unit (100 m)
  600 m3/h >> dry-pump
  2000 L/s >> TMP
  500 L/min >> TMP foreline pump
  1000 L/s >> IP
120123 VAC (YS)
DRY P
LCGT Vacuum System
pumping system (dry-pump and ion-pump)
2. System design

## Slide 9
control signal 0/1
pressure readout
120123 VAC (YS)
LCGT Vacuum System
vacuum system control
2. System design
arm tunnel, center/end room
BA gauge
ULVAC GI-M2
controller
PLC
Yokogawa
A/D
Yokogawa MW100
Micro IOC
Center control
EPICS
Pirani gauge
Convectron type
Gate valve
（actuator）
Ion pump
VARIAN type
Ion pump
Gamma type
Ethernet
RS485
→ status: 3
← control: 3
→ status: 4
← control: 2
→ status: 2
← control: 1
controller
controller
controller
controller
A gate valves of large diameter takes 40 seconds for closing.

## Slide 10
120123 VAC (YS)
ETMYA
ITMY
ETMXA
ITMX
BS
PRM
PR3
PR2
SRM
SR3
SR2
PD
MCF
MC
3000 m
iLCGT
iLCGT
MMT
ETMYB
ETMXB
bLCGT | HR Center |  | HR Chord Center |  | Center of Mass | 
 | x [m] | y [m] | x [m] | y [m] | x [m] | y [m]
PRM | -19.46123288 | 0.21377918 | -19.46120738 | 0.21377925 | -19.51120721 | 0.21364765
ETMX | 3.026507E+03 | -0.093909 | 3.026507E+03 | -0.093909 | 3.026582E+03 | -0.093909
ETMY | 0.060725 | 3.023222E+03 | 0.060725 | 3.023222E+03 | 0.060725 | 3.023297E+03
PR2 | -4.7006812 | 0.25262878 | -4.69784943 | 0.25267014 | -4.64785476 | 0.253400400
SRM | 0.13541083 | -19.41723194 | 0.13541104 | -19.41720645 | 0.13501125 | -19.46720485
SR2 | 0.25343357 | -4.65710099 | 0.25349012 | -4.65426948 | 0.25448854 | -4.60427945
SR3 | -0.10002163 | -15.71840442 | -0.10001528 | -15.71808665 | -0.1010137 | -15.76807668
PR3 | -15.7637214 | -0.04148993 | -15.76340359 | -0.04148529 | -15.81339825 | -0.04221555
BS | -7.63E-16 | -6.94E-18 | -7.63E-16 | -6.94E-18 | 0.02828427 | -0.02828427
ITMX | 26.50720079 | -0.09390876 | 26.50720439 | -0.09390876 | 26.43220439 | -0.09390876
ITMY | 0.06072497 | 23.22189738 | 0.06072497 | 23.22190098 | 0.06072497 | 23.14690098
iLCGT | HR Center |  | HR Chord Center |  | Center of Mass | 
 | x [m] | y [m] | x [m] | y [m] | x [m] | y [m]
PRM | -19.51677794 | 0.21417189 | -19.51675245 | 0.21417196 | -19.56675227 | 0.21404036
ETMX(iEOAX) | 3.001576E+03 | -0.093237 | 3.001576E+03 | -0.093237 | 3.001626E+03 | -0.093237
ETMY(iEOAY) | 0.061257 | 2.998291E+03 | 0.061257 | 2.998291E+03 | 0.061257 | 2.998341E+03
PR2 | -4.7281891 | 0.25309529 | -4.72535740 | 0.25313951 | -4.67536350 | 0.253920280
SRM | 0.13550922 | -19.47206837 | 0.13550943 | -19.47204288 | 0.13510964 | -19.52204128
SR2 | 0.25374881 | -4.68452802 | 0.25380268 | -4.68169646 | 0.25475388 | -4.63170551
SR3 | -0.07797132 | -15.7185301 | -0.07796549 | -15.71821231 | -0.07888163 | -15.76820392
PR3 | -15.76285711 | -0.06258281 | -15.76253931 | -0.06257763 | -15.81253268 | -0.06339185
BS | -4.09E-16 | 3.47E-18 | -4.09E-16 | 3.47E-18 | 0.02832301 | -0.02824548
ITMX(iIAOX) | 53.21470655 | -0.09323681 | 53.21471015 | -0.09323681 | 53.16471015 | -0.09323681
ITMY(iIAOY) | 0.06125676 | 49.92936491 | 0.06125676 | 49.92936851 | 0.06125676 | 49.87936851
less changes “i” to “b”
LCGT Vacuum System
layout
2. System design

## Slide 11
120123 VAC (YS)
ETMYA
ITMY
ETMXA
ITMX
BS
PRM
PR3
PR2
SRM
SR3
SR2
PD
MCF
MC
3000 m
iLCGT
iLCGT
MMT
ETMYB
ETMXB
i-b | HR Center |  | HR Chord Center |  | Center of Mass | 
 | delta x [mm] | delta y [mm] | delta x [mm] | delta y [mm] | delta x [mm] | delta y [mm]
PRM | -55.545 | 0.393 | -55.545 | 0.393 | -55.545 | 0.393
ETMX | -24931.300 | 0.672 | -24931.300 | 0.672 | -24956.300 | 0.672
ETMY | 0.532 | -24931.340 | 0.532 | -24931.340 | 0.532 | -24956.340
PR2 | -27.508 | 0.467 | -27.508 | 0.469 | -27.509 | 0.520
SRM | 0.098 | -54.836 | 0.098 | -54.836 | 0.098 | -54.836
SR2 | 0.315 | -27.427 | 0.313 | -27.427 | 0.265 | -27.426
SR3 | 22.050 | -0.126 | 22.050 | -0.126 | 22.132 | -0.127
PR3 | 0.864 | -21.093 | 0.864 | -21.092 | 0.866 | -21.176
BS | 0.000 | 0.000 | 0.000 | 0.000 | 0.039 | 0.039
ITMX | 26707.506 | 0.672 | 26707.506 | 0.672 | 26732.506 | 0.672
ITMY | 0.532 | 26707.468 | 0.532 | 26707.468 | 0.532 | 26732.468
less changes “i” to “b”
LCGT Vacuum System
layout
2. System design

## Slide 12
x
y
Yopt
horizontal plane at center room
horizontal plane at X end
Y arm
optical plane of interferometer
X arm
BS
ETMX
ETMY
X arm
Y arm
ITMX
ITMY
Xopt
Yopt
O
Xopt
O (BS)
iLCGT
1/300
120123 VAC (YS)
horizontal plane at Y end
LCGT Vacuum System
layout
2. System design
** floors for installing chambers; horizontal

## Slide 13
x
y
Yopt
horizontal plane at center room
horizontal plane at X end
Y arm
optical plane of interferometer
X arm
Xopt
O (BS)
120123 VAC (YS)
horizontal plane at Y end
LCGT Vacuum System
layout
2. System design
** unit vectors
in each coordinate
** transformation matrix for 4 sets of coordinates

## Slide 14
** transformation matrix for 4 sets of coordinates
x
y
Yopt
horizontal plane at center room
horizontal plane at X end
Y arm
optical plane of interferometer
X arm
Xopt
O (BS)
120123 VAC (YS)
horizontal plane at Y end
LCGT Vacuum System
layout
2. System design
** description
in other coordinate

## Slide 15
x
y
Yopt
horizontal plane at center room
horizontal plane at X end
Y arm
optical plane of interferometer
X arm
BS
ETMX
ETMY
X arm
Y arm
ITMX
ITMY
Xopt
Yopt
O
Xopt
O (BS)
iLCGT
1/300
120123 VAC (YS)
horizontal plane at Y end
LCGT Vacuum System
layout
2. System design
500 m
** a set of two 1.5-km long interferometer for geophysics

## Slide 16
Apr 2014
Apr 2014
Apr 2014 to Aug 2014
Sep 2014
Oct 2014 to Mar 2015
1) manufacturing 478 of tubes; from Apr 2011 to Mar 2013
2) manufacturing chambers; from Sep 2012 to Mar 2014
3) installing chambers in X end; Apr 2014
4) installing tubes in X arm; from Apr 2014 to Aug 2014
5) installing chambers in Center Room; Sep 2014
3) installing tubes in Y arm; from Oct 2014 to Mar 2015
5) installing chambers in Y end; Mar 2015
X arm pump down; Sep 2014
Y arm pump down; Mar 2015
120123 VAC (YS)
LCGT Vacuum System
expected schedule for installing (1)
3. Schedule
ETMYA
ITMY
ETMXA
ITMX
BS
PRM
PR3
PR2
SRM
SR3
SR2
PD
MCF
MC
3000 m
iLCGT
iLCGT
MMT
ETMYB
ETMXB
carrying tubes from MOZUMI pithead

## Slide 17
Apr 2014
Apr 2014
Apr 2014 to Aug 2014
Sep 2014
Oct 2014 to Mar 2015
1) manufacturing 478 of tubes; from Apr 2011 to Mar 2013
2) manufacturing chambers; from Sep 2012 to Mar 2014
3) installing chambers in X and Y ends; Apr 2014
4) carrying tubes and lay on the supports;
          from Apr 2014 to Jun 2014
5) jointing tubes and pump down in X and Y arm
          from Jul 2014 to Mar 2015
6) installing chambers in Center Room; Jul 2014
X arm pump down; Oct 2014
Y arm pump down; Mar 2015
120123 VAC (YS)
LCGT Vacuum System
expected schedule for installing (2)
3. Schedule
carrying tubes from ATOTSU pithead
ETMYA
ITMY
ETMXA
ITMX
BS
PRM
PR3
PR2
SRM
SR3
SR2
PD
MCF
MC
3000 m
iLCGT
iLCGT
MMT
ETMYB
ETMXB

## Slide 18
** recovery “without breaking arm-vacuum”
  BA gauge failure: grid contamination, filament breakdown
  feedthrough erosion; high voltage feedthrough of ion pump
  power supply/controller erosion; ion pump, turbo molecular pump
  durability for humidity is being examined (50°C-98%, 7 cycle of 8h-operation/98h-off).
** recovery “by re-pumping arm”
  window (view port) break down
  crack and erosion in bellow joint
  metal gasket erosion
** safety “by closing gate valves (large dia,)”
    electric-power shut down
    anomalous pressure rise
ETMX
ITMX
3000 m
expected pump-down scheme (a 3-km arm)
  to 1 Pa >> few days by dry-pump
  to 10-6 Pa >> 50 hours by TMP
  to 10-7 Pa >> 500 hours by IP
(based on the ougassing rate in test tubes)
120123 VAC (YS)
IP
TMP
FL P
DRY P
LCGT Vacuum System
vacuum component failure and leakage
4. Risk management

## Slide 19
** “surface passivation process” of stainless steel prior to installation is planned by applying electro- polishing, and then followed by pre-baking treatment.
** “outgassing rate” of the order of 10-8 Pa m3 m-2 s-1, or lower is performed in J-PARC.
Appendix: outgassing data measured for technical surface
Chart: pumping time [hours]
Cu (PR-EF lining, 2nd pump.): ID 560 X L 3000 | 5.2e-08
120123 VAC (YS)
LCGT Vacuum System