External Review of
Main Interferometer Subsystem
2012/4/19
Yoichi Aso

---
Overview: Definition and scope of the subsystem
MIF includes: Arm cavities, Recycling cavities, Michelson part
Assigned tasks
Parameter selection of the MIF
Length, ROC, etc
Optical layout
Design of the control scheme
LSC, ASC
Lock acquisition scheme
Commissioning
For some reason, detector characterization is a part of MIF
Maybe the Instrument Control will be too

---
Interface with other subsystems
Mirror
Reflectivity, ROC, Loss, other specs.
Suspension
Isolation requirements
Actuator design
Installation procedure
Suspension+AOS
Local sensors
Laser
Power
IOO
Intensity Noise
Frequency Stabilization Path
Input/Output mode matching
Detection ports
OMC
Green lock
AOS
Scattered light
Digital System
Servo model
Analog electronics
PD, QPD, I-Q demodulator, servo
filters, etc
Vacuum
Layout

---
Design Status
IFO params	semi-final	LSC principle	final
ASC principle	preliminary	Servo modeling	TBD
Green Lock Design	preliminary	Optical Layout	final
Frequency Servo	preliminary	Commissioning Plan	TBD
OMC	early stage	Analog Circuits	TBD

---
bKAGRA

---
Requirements
Operate the interferometer with the following noise level
Duty factor > 80% (Interferometer lock > 99%)

---
bKAGRA Interferometer
Signal Ports
DARM: AS_DC
CARM: REFL_1I
MICH: REFL_1Q
PRCL: POP_2I
SRCL: POP_1I
Y-arm
ITMY
PRM
REFL
PR2
POX
ITMX
ETMX
Laser Beam
From MC
BS
X-arm
PR3
SR2
AS_DC
OMC
SR3
SRM
Carrier
f1 (PM)
f2 (PM)
f3 (AM)

---
Almost Final Design
Interferometer Parameter
Arm Finesse	1550	Arm Length	3km	PRG	10
SRM Reflectivity	85%	Input Laser Power	51W	PRC Length	66.6m
SRC Detuning	3.5°	SRC Length	66.6m	MICH Asym.	3.33m
g1	-0.786	g2	-0.602	g1・g2	0.473
R1	1.68km	R2	1.87km
For details:
http://gwwiki.icrr.u-tokyo.ac.jp/JGWwiki/LCGT/subgroup/ifo/MIF/OptParam
MIF Design Document: JGW-T1200913

---
g-factor selection
Beam Size
ITM: 3.5cm, ETM:4.0cm
4.0cm is the maximum spot size for 22cm mirror.
3.5cm beam at the central part is easier to handle.
Since ITMs have less coating layers, the thermal noise is not
compromised by 3.5cm beam size on ITMs.
Decided:
size 32 {g_1 cdot g_2 = 0.473}
Candidates
g1=0.786, g2=0.602 (R1=14km, R2=7.5km)
g1=-0.786, g2=-0.602 (R1=1.68km, R2=1.87km)

---
Arm Cavity HOM Resonances
No significant difference between Positive and Negative
See JGW-G1200789
Parametric Instability
Colored area = Unstable mode exists
by K. Shibata
See JGW-T1200787

---
Sidles-Sigg Instability (ASC)
WFS UFG for Pitch has to be lower than 1Hz
For arm power > 330kW, Pitch instability frequency > 1Hz
for positive g-factor
Negative g-factor never makes Pitch unstable
Current arm power (compromised to deal with high sapphire
absorption) = 250kW --> even positive g-factor is OK
Original (desirable) arm power = 376kW
        --> positive g-factor is not an option
Conclusion
Negative g-factor is better
Positive g-factor limits us to operate only with low power

---
TM Pitch Transfer Function

---
Gouy Phase Change in PRC/SRC
Factors to Consider
Higher Order Mode Resonances in RCs
WFS Signal
Stable Recycling Cavity → Suppression of TEM10/01
        → Small Arm WFS Signal
Smaller Gouy phase is better for WFS ? → Not So Simple
For PRC Gouy phase ~ 90゜is better for TEM10/01 transmission
WFS signal separation for auxiliary mirrors has to be checked
Large Gouy phase change requires small beam spot on PRM

---
HOM Power in PRC

---
SRC Carrier HOM Power Gain

---
Length Sensing and Control
Canonical IFO DOFs
(DARM, CARM, MICH, PRCL, SRCL)
Mirror displacement
size 20 {vec n_{nitalic s} }
Sensing Noise
Calibrated
Output
Displacement
noise
Detector
Sensing
Calibration
+
D
S
C
size 20 {vec x_{nitalic c} }
size 20 {vec n_{nitalic d} }
+
+
-
Actuator
Filter
F
A
+
F'
Feed forward
size 26 {vec x_{nitalic c} = (I+G)^{-1} cdot S cdot vec n_{nitalic s} + (I+G)^{-1} cdot SD cdot vec n_{nitalic d}}
size 26 {G equiv S cdot D cdot A cdot (F+F')}
D: Optical Gain Transfer Function calculated by Optickle
A: Mechanical TFs of the mirror suspensions + optical spring stiffness
F: Feedback filters
F': Feed forward matrix
S: Sensing matrix (here I just used identity matrix)

---
bKAGRA Interferometer
Signal Ports
DARM: AS_DC
CARM: REFL_1I
MICH: REFL_1Q
PRCL: POP_2I
SRCL: POP_1I
Y-arm
ITMY
PRM
REFL
PR2
POX
ITMX
ETMX
Laser Beam
From MC
BS
X-arm
PR3
SR2
AS_DC
OMC
SR3
SRM
Carrier
f1 (PM)
f2 (PM)
f3 (AM)

---
LSC Loop Noise Couplings
UGF DARM: 200Hz, CARM: 10kHz, Others: 50Hz
Feed Forward Gain: 100

---
ASC
by Y. Michimura
WFS+OpLev

---
Total Angular Noise Couplings to DARM
Pitch
Yaw

---
Practical Noise Requirements
Auxiliary Mirror Displacement Noise
Laser Frequency Noise / Intensity Noise
RF Oscillator Phase Noise / Amplitude Noise
Scattered Light Noise Couplings
All the requirements include a safety factor of 10

---
Displacement Noise Requirements to Auxiliary Mirrors

---
Displacement Noise Requirements to Auxiliary Mirrors
CLIO Thermal Noise (room temp.)

---
Frequency Stabilization Servo
MC
CARM
Laser
Reference
Cavity

---
Frequency Stabilization Servo
Black Box
MC
CARM
Laser
Reference
Cavity
UGF=10kHz

---
Laser Frequency Noise Requirements
Reference Cavity Stability

---
In terms of the MC displacement noise
Reference Cavity Stability
MC Suspension Noise

---
Laser Intensity Noise Requirements
Relative Intensity Noise (RIN)
RIN = 10-9 is tough to achieve

---
RF Oscillator Phase Noise Requirements
Single Sideband (SSB) Phase Noise =
(Phase Noise SB Power)/(Carrier Power)
Wenzel OCXO

---
RF Oscillator Amplitude Noise Requirements
RF Amplitude 1V - Noise 100nV

---
Why are noise requirements for DRSE more stringent ?
Carrier
Detuning SRC
f1 sidebands
Pure Phase Modulation
Mixed Amplitude/Phase
Modulation
Offset in the error signals

---
Solution ?
Detuning SRC
Prepare f1 as a mixture
of AM/PM
Pure Phase Modulation

---
Scattered Light Modeling
Scattering Surface
a
To DARM
Es
Es
ETM
size 28 {}
Scattered Light Noise

---
Requirements for Scattered Light Noise
size 38 {E_"s" cdot a " [" sqrt{"W"} cdot "m"/sqrt{"Hz"} "]"}
Actual Scattered Light Noise: Calculation by Mike Smith
Proper arrangement of baffles will reduce the scattered
light noise below the requirements

---
Optical Layout Issues
Smaller Wedge Angles
New POX/POY Extraction Ports
RC Mirror ROC Error Tolerance
Wedge Angle Error Tolerance

---
New Layout Philosophy
A Big Help from Mike Smith !
Before
POY
POX

---
New Layout Philosophy
A Big Help from Mike Smith !
After
POY
POX

---
ITMY
POY
ITMX
POX

---
However, we cannot use Mike's layout as is.
New functionalities added to GTrace
Non-sequential trace mode
Automatically trace all stray beams
Termination criteria: power, number of AR reflections
Better rendering of beam width

---
Wedge Angle Error Tolerance
ITM Wedge Angle = 0.013(+0.003/-0.005) deg
BS Wedge Angle = 0.07(+0.005/-0.01) deg
Wedge Angle Error → Optimal incident angle to BS from PRC changes
→ Positions of PR3/SR3 and ITMs change < a few mm
POX/POY beam separation from the main beam changes
Scan all possible combinations of error polarities in BS, ITMX and ITMY
     → Worst case scenario

---
POY
POX

---
ROC Error Tolerance
ROC Error of PR2/PR3 → Mode Mismatch , Gouy Phase Error
ROC error can be corrected by changing Lp2
PRM
PR2
PR3
SR3
SR2
SRM
ITMY
ETMY
ITMX
ETMX
Larm
Larm
Lp1
Lp2
Lp3
Ls1
Ls2
Ls3
Lmx
Lmy
BS
Lmi = (Lmx + Lmy)/2
Las = Lmy - Lmx
Lprc = Lp1+Lp2+Lp3+Lmi
Lsrc = Ls1+Ls2+Ls3+Lmi
From MC

---
ROC error correction by Lp2
ROC Error = 1cm
Slope=0.76deg/mm

---
Green Lock Setup
D. Tatsumi & K. Arai

---
Arm Cavity Stability after Green Lock
Requirement: 0.3pm RMS

---
Arm Cavity Stability after Green Lock
Requirement: 0.3pm RMS

---
OMC
Conceptual Design by K. Somiya
Transmission Power after OMC
Signal Loss/Shot Noise Increase
by TM ROC Errors

---
Prototype Test
ISC simulation using virtual interferometer on RTS
Green injection table will be tested beforehand
(Kashiwa ? NAOJ ?)
OMC
AM/PM mixed modulation
Quality Assurance
Commissioning is QA for MIF

---
Risk Management
Risk factors
Arm cavity loss is too large: PRC under-coupled
=> Prepare a spare PRM with lower reflectivity
Arm cavity reflectivities are so matched that the homodyne phase
cannot be set to around 45°
=> Forget about BAE
SRC detuning makes the practical noise couplings bigger
=> AM/PM mixture for f1 ?
Commissioning takes too long (very likely to happen)
=> Work even harder folks !
Lack of man power
 => Get money to hire people

---
iKAGRA

---
iKAGRA Design
Target Specs
Stably operate Fabry-Perot Michelson Interferometer
No sensitivity requirement
Final(?) Design
Based on bKAGRA design
Use Silica TMs
ROCs are the same as bKAGRA
Locations of TMs are shifted by 26.6m/30m from bKAGRA
ITMs laterally shifted by 7cm
Arm GV is shifted accordingly
No folding part optimization
PRM will not be installed (PR2, PR3 will be)
Tweak the Input MMT to get a reasonable mode matching
SRC is isolated vacuum wise.
Pick-off the AS beam before the GV
Green lock may be tested with X-arm

---
(Virtual) Layout Change in iKAGRA
Location of ITMs: Shifted by 26.6m from the bKAGRA position
Location of ETMs: Shifted by 30m from the bKAGRA position
ITM lateral shift: 7mm
PR2
Move
Move
lateral shift
7mm
PRM
26.6m
30m
ITMX
ETMX
BS
PR3

---
Diagonal incidence to compensate
for the wedge
ITM
(bKAGRA)
Arm Cavity Axis
From BS
From BS

---
Just moving the ITM 26.6m ahead will
mis-center the beam
ITM
(iKAGRA)
ITM
(bKAGRA)
From BS
26.6m

---
Mode Matching
Without PRM
With PRM
DRMI
JGW-G1200763-v1

---
iLIGO Test Masses for iKAGRA ?
Making new silica TMs costs a lot of money (~$1M)
Can we use the iLIGO TMs as is for iKAGRA ?
No ! the wedge angle is too much different
However
Is this possible ?
Re-polish and re-coat the AR surface
Even so, we have to at least make a new IMMT
Does this save money ?
Seems like cost more time.

---
Vacuum Isolation of SRC
Separate Vacuum
during iKAGRA
Pick-off the AS
beam here.

---
Quality assurance
Each component will be tested by corresponding subsystem
(e.g. mirrors, laser, etc)
The quality of MIF is the noise level of the interferometer
It is guaranteed that the quality will not meet the requirement
at first
Thorough and careful planning for the commission is the best
quality assurance

---
Risk Management
Risk factors
StackB is too noisy
=> Strong damping, strong actuators at the expense of noise
Laser frequency noise is too large
=> Spend more time on the frequency stabilization servo
Mirror quality is poor
=> ....
Commissioning takes too long (very likely to happen)
=> Work harder folks !

---
Schedule
Task list during tunnel excavation
(to be finished by March 2014)
LSC/ASC Servo model including hierarchical actuation
Frequency servo topology
Servo model noise analysis
Servo model implementation on RTS
Green lock detailed design
Green injection system assembly
Commissioning planning
Suspension positioning method
Optical bench layout (with IOO)
Analog electronics, design, manufacture, test

---
Installation & Commissioning Scenario
iKAGRA

---
Schedule & Installation Scenario
bKAGRA

---
Human Resources
What we are hoping for ...
2012 - 2013 (Tunnel excavation period)
Fully engaged post-doc: 2
Grad. students: 4
Electric engineer: 2
2014 - 2018 (Commissioning period)
Fully engaged post-doc: 5
Grad. students: 10
Electric engineer: 2
What we have now ...
Me, Kentaro, some students, working part time
on KAGRA MIF
Most staff scientists are chiefs of other subsystems
Many people pledged to spend ~ 5% of their time
→ Sorry, they don't count

---
Instrument Control
What we want to do
Remote monitoring of instruments deployed across the site
Vacuum equipments, Cryocoolers, Lasers, etc ...
Remote control of instruments
Automatic and quick response to accidents (Interlock)
Closing gate valves in case of vacuum break
Shutting down lasers to prevent harm to human
Scope
All the instruments not controlled by the RTS
Vacuum system
Cryogenic system
Laser
Environmental monitors (?)

---
Implementation Plan
Control Room
ICM
Instrument 1
EPICS
EPICS
TCP/IP
Network
EPICS
EPICS
ICM
Instrument 2
EPICS
EPICS
ICM
Instrument 3
EPICS
Frame Builder
Instrument Control Module(ICM)
PLC
EPICS
IOC
PLC: Programmable Logic Controller
IOC: Input Output Controller
Frame Files

---
Current Status
We don't have man power !
Outsource to companies
Talking with two companies, big and small
Costs money
We don't have money !
We need to get quotes from the companies
They need rough number of channels