## Slide 1: Reserve

## Slide 2
2
2
Motivation : Performance test of cryogenic payload
Initial cooling test with liquid nitrogen (Jecc Torisha)
: 30th Aug. 2013
Delivery : 27th of Sep. 2013
1/4 Crystal

## Slide 3: Vibration measurement of KAGRA cryostat
Preliminary
Vibration from the radiation shield through heat links.
Recoupling into main laser via scattered light
Noise of the interferometer
Measure the vibration of the radiation shield, and estimate the influence on the sensitivity of KAGRA.
Heat links
8K
80K
Laser
Radiation shield
Cooling bar
Cryo cooler
Cryo cooler
mirror
Baffle
Measurement at 10K
Consistent
Increase because of the cryostat structure
The impact on the sensitivity of KAGRA
(Horizontal component)
> 100 times
T=10K
The noise from the horizontal vibration component is significantly lower than the design sensitivity.
Notes: Analysis is on going now by Chen Dan kun.

## Slide 4
Main Interferometer Subsystem
Basic Optical design has been fixed
Phys. Rev. D, 88, 043007 (2013)
Detailed design
Actual servo topology, filter design, noise analysis
Analog electronics
Optical layout including auxiliary beams
Output optics design needs more work
OMC design
Output Faraday
Suspensions for output optics (OMC, Faraday, etc)
Notes: Basic optical design of the main interferometer was fixed a while ago. We recently published a paper on this.
Now, the main focuses of the main interferometer group are the detailed design of the actual servo system, developing analog electronics for high frequency servos, detailed optical layout including auxiliary beams and so on.
However, there remain an area more work is necessary, that is, the output optics system.
The design of the OMC is still immature and the design of the other optics for the output path are not yet finished. So we are working hard on these.

## Slide 6: Black coating on baffle surface
The black coating on the baffles is classified into a few difficulty levels.
Our default candidate was a diamond-like carbon (DLC) coating, but we found it was not so black (~20% reflectivity) at 1 um. We look for other candidates, even though the DLC is the best choice if only the vacuum compatibility is considered.

## Slide 7: Reflectivity of candidates
Solblack, a kind of electroless Ni-P plating by Asahi Precision Co., has the lowest reflectivity!!
Notes: P; phosphorus [fάsfərəs | fɔ́s-]｟記号P｠
ECB: Electric Chemical Buffling

## Slide 8: Reflectivity of candidates
in ATC-NAOJ
Notes: ECB: Electric Chemical Buffling

## Slide 9: Scattering distributions
SolBlack on SUS BA
Solblack on SUS 2B
Solblack on A1050 ,ECB
ECB (Electro-Chemical Buffing)
Scattering probability distribution per unit solid angle
Scattering profiles differ as the base material’s surfaces!
Note: Different from “BRDF” .
FWHM :About 4 deg
Our design principle: use mirror-like surfaces and suppress wide-spread scatter.
Candidate
Rotating detector
Source: 1um
Goniophotometer
Measurement by Y. Niwa in JASMINE Project, NAOJ
Measurement system also belongs to their project.

## Slide 10: Other tests
Tests for the vacuum compatibility is done (after baking in vacuum).
Tests for cryogenic compatibility is almost done.
Cryogenic Test in 70K (for the surface of 5-m cryoduct shield)
First, the reflectivities of a sample in 1um (YAG) and 10um (CO2 laser) are measured.
Next, the sample is cooled in liquid nitrogen.
And then, a visual inspection is perfomed.
Again, the reflectivities of the sample in 1um and 10um are measured. (cryo-resistant test)
Liquid N2
Sample: SolBlack on Al
by Asahi-Precision
More other test: Reflectivity check during cooling.
Cryostat
CO2 laser (10um)
The same test w/ 1um laser was done.
Remaining test: compatibility for 8K environment (now ongoing...)

## Slide 11: Ultimate design sensitivity of KAGRA
This sensitivity achieves 167 Mpc for
coalescence GW of 1.4 Ms NSB.
Duty factor: > 80%
Expected event rate
20kg test mass at 20 K
Suspension of Q=108
Bulk Q of mirror 108
400kW cavity power
Q of coating: 104
DC read out
9.8
yr-1
-6.6
+14

## Slide 12: Interferometer configuration
Input/Output Optics
- Beam Cleaning and stab.
- Modulator, Isolator
- Fixed pre-mode cleaner
- Suspended mode cleaner
Length 26 m, Finesse 500
- Output MC
- Photo detector
Main Interferometer
- 3 km arm cavities
- RSE with power recycling
- Cryogenic test masses
Sapphire, 20K
‘Type-A’ vibration isolator
Cryostat + Cryo-cooler
- Room-temp. Core optics
(BS, PRM, SEM, …)
Laser Source
- Wavelength 1064 nm
- Output power 180 W
High-power MOPA
ETM
ITM
Y-arm cavity
X-arm cavity
Length 3,000 m
Finesse 1,550
ITM
ETM
BS
SEM
PRM
26-m MC
Power-recycling
Gain ~11
Laser
825 W
RSE: (Resonant
sideband Extraction)
Signal-band Gain ~15
Detuned RSE
(Variable tuning)
Power ~400 kW
Input
Bench
Power
~180 W
80 W

## Slide 13
Fabry-Perot Michelson
Power Recycled Fabry-Perot Michelson with broadband RSE
Cryogenic mirror, output MC
Without cryogenic & PR
Fused silica mirror
Operational run in 2015
Operation starts in 2017
Observation in 2018
From iKAGRA to bKAGRA

## Slide 14: Cryogenics : Cryostat assembling and performance test
Assembling and performance tests of Cryostats are in progress in the TOSHIBA Keihin Product Operations.
Cooling test of Cryostat No.1 was completed.
  It took about 2 weeks of cooling without artificial heat load.
  Rate of radiation cooling of Al sphere was agreed well with a result of simulation.
Specific test items
  Cryostat No.1 : Cooling test of 105 Al sphere without DLC. Done.
  Cryostat No.2 : Cooling test of 105 Al sphere with DLC and vertical vibration measurement of 8K shield by accelerometer (Roma Univ.).
  Cryostat No.3 : Cooling test of 1/2 dummy payload with DLC coating and sapphire dummy mirror and horizontal vibration measurement of 8K shield by accelerometer (ICRR) .
  Cryostat No.4 : Cooling test of 105 Al sphere with Solblack coating.
Cryostat No.1
1/2 scaled payload in the cryostat No.3.
Accelerometer and black Al sphere.
Installation of black Al sphere

## Slide 15
19 standard filters have been built and delivered: final assembly of the GAS filters is ongoing
Six out of 11 pre-isolator top filters have been finished
Prototypes of a complete chain pre-isolator, standard filter, bottom filter and optical payload have been completed and are under tests in Japan
Vibration Isolation
Main SAS: 13m tall
2 stages of floor

## Slide 16: SAS in digital control
Prototype test of Pre-isolator in Kashiwa
Digital control of Inverted Pendulum for horizontal isolation 3 degrees of freedom using LVDT&voice coil (sensors of acceleration) was successively conducted.
Pre-isolator cosists of three set
of Inverted pendulums and
vertical filter
Digital control system is constructed in collaboration with aLIGO

## Slide 17: Core optics
iKAGRA cavity mirrors
ETMs: iLIGO mirrors
ITMs: reformed iLIGO mirrors (wedge angle must be changed)
BS:
S-Polarization
37cm diameter , 8cm thickness
iKAGRA cavity mirrors
22cm diameter , 15cm thickness
C-Axis Sapphire
Recycling & Signal extraction cavities mirrors
reformed iLIGO mirrors
Radius of curvature must be fabricated with extremely small uncertainties.
Ex. PR2: R=(-3.0764±0.01) m PR3: R=(24.9165±0.01) m
MC &MT mirrors
Polish of MC mirrors has been inspected at LMA.
Flat Mirror： < 1 Årms
~150μm
Large crystal

## Slide 18
Laser : R&D for high power
The highest power was 78 W obtained from two 41-W outputs.
The performance of each component is being tested.
Total system assemble will be started soon as the next step.
Wave-front distortion caused by a solid-state amplifier
Coherent addition
Bright port 78W
Dark port 4W
The wave front distortion was corrected by using a deformable mirror.