## Slide 1: Prototype Test of Vibration Isolation System Current Status & Schedule
ICRR, D2
Takanori Sekiguchi

## Slide 2: Prototype Test of SAS
Standard GAS Filter @Nikhef & Kashiwa (Jan. 2011~)
Pre-Isolator @Kashiwa (Aug. 2011~)
Payload @NAOJ (Aug. 2013~)

## Slide 3: Current Status

## Slide 4: Standard GAS Filter Test
Done at NIKHEF:
Tuning of resonant frequencies
Measurement of mechanical transfer function
Measurement should be done:
Tuning of magic wands (establish the procedure at Kashiwa and apply it to actual equipment in Akeno)

## Slide 5: Pre-Isolator Prototype Tests
Done at Kashiwa:
Tuning of GAS & IP resonant frequencies
Digital (standalone) system setup
IP control in 3 DoFs with LVDT and geophone
DC control of IP with stepper motors
Test should be done:
Torsion mode damping

## Slide 6: Problems in Pre-isolator Prototypes
Easily broken stepper motor drivers (circuits used in TAMA, manufactured >10 years ago)
 New driver should be designed soon…
Low frequency (<0.1 Hz) noise of geophone preamp circuits (thermo-electromotive effect coupled with temperature fluctuation)
 Need temperature stabilization with an envelope

## Slide 7: Vacuum Envelope for Geophone
Preamp and geophone are equipped inside

## Slide 8: Payload Prototype Tests
Done at NAOJ:
Assembly test of TM suspension
Order components for OSEM driver
Test should be done:
OSEM calibration
Assembly of RM, IM, IRM suspension
Mechanical TF check
Active control with OSEM

## Slide 9: Test Plan for Full Type-B Prototype

## Slide 10: Important Schedule
Pre-isolator and its control system will be delivered to NAOJ this month. (on 10/22-23)
Installation of the actual vibration isolation system for i-KAGRA will start in the middle of FSY 2014 (~ Oct. 2014).
Now
Installation
Start
i-KAGRA
operation
~9 months left

## Slide 11: What We Learned from TAMA-SAS ??
Angular fluctuation of the mirror was more critical in TAMA.
  Need tools to adjust DC alignment
  Need dampers for torsion modes
& What is New from TAMA-SAS??
Actively controlled stage with OSEM. (c.f. eddy current damping in TAMA-SAS)
We should spend time for testing these items.
Picomotor
For yaw alignment
Motor + moving mass
For pitch alignment
OSEM
Torsion damper

## Slide 12: Main Action Item for Full Type-B Test
Measure angular fluctuation of the mirror with an optical lever.
  Check DC alignment adjustment
  Check efficiency of torsion mode damping
  Establish control scheme with OSEMs
Length sensing: Michelson interferometer from ground?? (TBD)

## Slide 13: Action Items (1)
Payload Test @20m Lab, NAOJ
  Tuning bottom GAS filter
  OSEM test & calibration
  Check motors for mirror alignment
  Check mechanical transfer functions with OSEMs
  Control test with OSEMs
Optical Lever Setup @TAMA
  Designing holders and covers for light source & detector
  Check commercial electric circuit for PSD
  Measure noise level

## Slide 14: Steps for Type-B Test
1. Pre-isolator Test in Vacuum
2. Full Type-B with Fixed Payload
3. Full Type-B Test
Payload Prototype
Optical Lever

## Slide 15: Action Items (2)
Pre-isolator Test in Vacuum @TAMA
  Re-setup of digital control system
  Compared with the result of in-air experiment
Full Type-B with Fixed Payload @TAMA
  Torsion mode damping test
  Check cabling effects
  Optical lever working test
Full Type-B Test @TAMA
  Mechanical TF check with equipped sensors and actuators
  Apply full control, check DC alignment adjustment

## Slide 16: Schedule
2013: payload test + sensor preparation
2014: Type-B test utilizing TAMA facility
Agatsuma & Joris join from NIKHEF

## Slide 17: END

## Slide 18: Appendix

## Slide 20: Problems in Pre-isolator Prototypes
Excess control noise at low frequency (<0.1 Hz) due to electronic noise in geophone preamp circuits
Possible source: thermoelectromotive effect coupled with temperature fluctuation of the circuit/geophone
Ideal Noise Level
Excess Noise