## Slide 1: OSEM LED Selection Measurements: Intensity Noise and Angular Distribution
NAOJ VIS Meeting, April 29nd, 2014

## Slide 2: Results from OSEMs last November (1/2)
LIGO papers have shown that OP232 LED intensity noise varies from batch to batch

## Slide 3: Results from OSEMs last November (1/2)
In addition, LIGO papers show different angular distribution and peak locations for different LEDs

## Slide 4: Measurement set-up
Determined with the first (!) fully assembled
OSEM what current to run the LED, now with collimator, to have -12V on UTIA output:
Rs,LED = 300W
ULED = 8V
ILED ≈ 25mA
For fast comparative measurement, LED is run in a simple set-up at 25 mA and ASD on the UTIA signal is taken
First, a dependence of intensity noise on the LED intensity magnitude is investigated

## Slide 5: Intensity noise LED current dependence

## Slide 6: Intensity noise first batch of LEDs

## Slide 7: Comparison between batches

## Slide 8: Measurement set-up
OSEM optics
Replace for trans- lation stage
LED
Lens (f = 8mm)
CCD Camera