## Slide 1: Error Contribution BRT
Junko Katayama
NAOJ / Ochanomizu University
2014/5/21

## Slide 2
TMS design
BRT
GPT
ETM
QPD
L0
L3
L4
L5
L1
L2

## Slide 3: L0,L1 Position Error Contribution to Gouy Phase difference
Error < ±1 [mm] required
OK

## Slide 4: L2,L3 Position Error Contribution to Gouy Phase difference
OK
Error < ±5 [mm] required

## Slide 5: L4,L5 Position Error Contribution to Gouy Phase difference
OK
OK

## Slide 6: RoC error 1
ex1) RoC of BRTlens1 front side
Error [％] | ω0 between QPD1 & QPD2 [m] | q parameter on GPT lens2 [m] | Compensation of Path Length between BRT lens1 & 2 [m] | Compensated ω0 with Gouy phase difference 90 degree [m]
+1 | No waist | 0.624248036004+0.0318213058677i | 1.822026123355 | 0.00016991
0 | 0.000168293 | -0.332940569997+0.083625468523i | 1.8022 | 0.000168293
-1 | No waist | 0.0943883935588+0.00627511054543i | 1.782374597325 | 0.000166672

## Slide 7: RoC error 2
ex2) RoC of BRTlens2 front side
Error [％] | ω0 between QPD1 & QPD2 [m] | q parameter on GPT lens2 [m] | Compensation of Path Length between BRT lens1 & 2 [m] | Compensated ω0 with Gouy phase difference 90 degree [m]
+1 | 0.000136795 | -0.228402214777+0.0552519530522i | 1.8005001057955112 | 0.000166705
0 | 0.000168293 | -0.332940569997+0.083625468523i | 1.8022 | 0.000168293
-1 | 0.000219617 | -0.500371551092+0.142409759343i | 1.8039005533490906 | 0.000169909

## Slide 8: RoC error 2
Error [％] | ω0 between QPD1 & QPD2 [m] | q parameter on GPT lens2 [m] | Compensation of Path Length between BRT lens2 & GPT lens1 [m] | Compensated ω0 with Gouy phase difference 90 degree [m]
+1 | 0.000136795 | -0.228402214777+0.0552519530522i | -1.17918871025 | No waist
0 | 0.000168293 | -0.332940569997+0.083625468523i | 1.3 | 0.000168293
-1 | 0.000219617 | -0.500371551092+0.142409759343i | 6.12412690512 | 0.000272185

## Slide 9: Conclusion
1. L1 between BRT lenses & L3 between GPT lenses are sensitive
(Errors < ±1mm Required)
2. RoC errors can be compensated by L1