## Slide 1: Monolithic Accelerometers on the Inverted Pendulum: status report
NAOJ GWPO Meeting, April 21st, 2016

## Slide 2: Reducing the velocity starts at the top
Damping of the modes in the
chain up until the payload is done
at the top stage
Modes give rise to more rms velocity, so have to be damped
The position towards the ground has to be controlled, without introducing the mircroseism

## Slide 3: The Inverted Pendulum stage control
= LVDT
= Geophone/ Accelerometer
LVDTs and Seismometers (here geophones) are in ‘pin-wheel’ configuration and signals are in geometrically reconstructed
to x,y and yaw

## Slide 4: LVDTs also ‘measure’ ground motion
Ground
Most of the ground velocity rms comes from the microseism
Measuring the ground and subtracting it, is not foreseen in KAGRA
so one has to find another way

## Slide 5: LVDT and Seismometer are blended
For example, for 0.1 Hz LVDT blending filter still has > 0.1 magnitude at the microseism
Lowering blending frequency can only be done if the seismometer is good enough

## Slide 6: Seismometer noise can spoil performance
LVDT
Geophone
Different seismometer?
Strange behavior 0.15-0.3 Hz

## Slide 7: The LVDT Monolithic Accelerometer
Because of inverted pendulum, natural frequency is low
Displacement of the mass is read out by LVDT
A coil-magnet actuator is used to calibrate the sensor

## Slide 8: Accelerometer f0 and Q determination

## Slide 9: Putting the parts together…

## Slide 10: 2 LVDTs were found damaged…
… and an attempt was made to repair them in situ
But, in the end, LVDT 1 and LVDT 3 were not working  Shipped all to company 2016.04.08

## Slide 11: Determining the V/mm conversion
a
ga
Alessandro calibrated all accelerometers two years ago:
All had GLVDT ~ 60 V/mm
INA103 amp:
Settings now at KAGRA give GLVDT ~ 110 V/mm
This has to be verified by tilt calibration

## Slide 12: TF measurement of unloaded IP
L/T
Y

## Slide 13: Loading the IP with tuning masses
On 2016.04.08, Takahashi-san and Fujii-kun loaded the IP:
Adding 230 kg gave fIP= 540 mHz
Takanori had about 720 kg for fIP= 540 mHz and about 1070 kg for fIP= 80 mHz
Current fIP status

## Slide 14: Accelerometers need air shields
Clean booth air flow

## Slide 15: Accelerometers vs. Geophones

## Slide 16: Blending lower with GEOs and ACCs
Blending at ~20 mHz already seems to be enough

## Slide 17: I can’t measure remotely
Luckily, there’s
can help on site
Thanks everyone!

## Slide 18: After iKAGRA, let’s get a working set-up!
Congratulations on the progress on iKAGRA!
After 2nd observation run: - all LVDTs are at a company for repair and return in May
- IP is loaded with tuning masses
- tilt calibration still has to be done to verify the 110 V/mm
Meanwhile, I’ll try to simulate more the expected performance of the accelerometers
when used in-loop to see if some tweaking (e.g. gain) is necessary
Maybe LVDTs could be less noisy? Now above 10 nm/rtHz below 1 Hz

## Slide 19: ADC noise model

## Slide 20: Mass thermal noise in air