## Slide 1
The BF dampers use common coils in the LVDT and the actuator. The actuator signals coupled to the LVDT signals in O3. The LVDT drivers were replaced to the modified ones. The BF damper may suppress the lowest mode of filter chain at 0.15Hz.
Plan1-1. Improvement of LVDT driver for BF damper

## Slide 2
Improvement of LVDT drivers for BF damper
Measured transfer functions from an actuator to the LVDT readout of the same channel. There is a dramatic improvement in eliminating the spurious coupling at high frequencies.
H1
H2
H3
V1
V2
V3
https://klog.icrr.u-tokyo.ac.jp/osl/?r=16534

## Slide 3
The servo type accelerometers (IX, IY) and the commercial geophones (EX, EY) are used as inertial sensors in Type-A towers.
The sensitivity of present inertial sensors are not good enough to control the IPs around 0.1Hz.
Plan1-2. Improvement of inertial sensors for IP control
We replace the inertial sensors to better accelerometers.
  Replace the present position sensors.
  Tune the folded pendulums to 0.2Hz.

## Slide 4
Some kinds of position sensors (photointerrupter, fiber LED, white light interferogram, and LVDT) have been evaluated. The LVDT had the best sensitivity around 0.1Hz.
Prototype accelerometer with the LVDT was tested at the KAGRA site.
Sensitivities of some kinds of sensors
(Geophone)
Measured spectra at the KAGRA site
Prototype test at the KAGRA site
One order improvement at 0.1Hz is expected.

## Slide 5
The F0 had overloaded by 4-10kg. The F0 keystones were fixed in O3. Three (of six) blades (for 297kg load) were replaced to strong blades (for 347kg load). The F0 can work as a vertical filter and the height of keystone became adjustable.
Plan3-1: Replacement of GAS blades for F0 in IX/IY
The F0 was tuned to 0.2Hz (as single pendulum). The adjustable range was extended to more than +/- 3.5mm.
Plan3-2. Extension of adjustable range for F0 in IX/IY